## Slide 1: Medical Device Cybersecurity: FDA Perspective Suzanne B. Schwartz, MD, MBA Associate Director for Science & Strategic Partnerships Office of the Center Director Center for devices & Radiological Health October 19, 2016

## Slide 2: Agenda
Background
  Cybersecurity Landscape in HPH Sector
  Presidential Executive Orders and National Institute of Standards and Technology (NIST) Framework
  CDRH/FDA Cybersecurity Activities
Total Product Life Cycle (TPLC) Framework
  Premarket & Postmarket Cybersecurity Approach
Next Steps

## Slide 3: Framing The Issue: Environment
The health care and public health (HPH) critical infrastructure sector represents a significantly large attack surface for national security today
  Intrusions and breaches occur through weaknesses in the system architecture
Connected medical devices, like all other computer systems, incorporate software that are vulnerable to threats
We are aware of cybersecurity vulnerabilities and incidents that could directly impact medical devices or hospital network operations
When medical device vulnerabilities are not addressed and remediated, they can serve as access points for entry into hospital/health care facility networks
  May lead to compromise of data confidentiality, integrity, and availability
  May be a safety issue

## Slide 4
Contain configurable embedded computer systems
Increasingly interconnected
Wirelessly connected
Legacy devices
Varied responsibilities for purchase, installation and maintenance of medical devices, often silo-ed
Variable control over what is placed on the network
Inconsistent training and education on security risks
4
Medical Device Cybersecurity
MEDICAL DEVICES
USE ENVIRONMENT

## Slide 5
Network-connected medical devices infected or disabled by malware
Malware on hospital computers, smartphones/tablets, and other wireless mobile devices used to access patient data, monitoring systems, and implanted patient devices
Uncontrolled distribution of passwords
Failure to provide timely security software updates and patches
Security vulnerabilities in off-the-shelf software designed to prevent unauthorized device or network access
Examples of Observed Medical Device
Cybersecurity Vulnerabilities
Privileged and Confidential

## Slide 6: Executive Orders (EO), Presidential Policy Directives, and Framework to Strengthen Critical Infrastructure Cybersecurity
EO 13636 (Feb 2013)
“We can achieve these goals through a partnership with the owners and operators of critical infrastructure to improve cybersecurity information sharing and collaboratively develop and implement risk-based standards.”
PPD 21 (Feb 2013)
NIST Voluntary Framework (Feb 2014)
EO 13691 (Feb 2015) – establishment of Information Sharing and Analysis Organizations (ISAO)

## Slide 7
CDRH/FDA Cybersecurity Activities
FDA Safety Communication
Draft Premarket Guidance
Begin Coordination with DHS
Recognize Standards
Establish Incident Response Team
Final Premarket Guidance
MOU with NH-ISAC
Public Workshop
Product-specific safety comm
Build ecosystem/collaboration
2013
2014
2015
2016
Draft Postmarket Guidance
Public Workshop

## Slide 8: Premarket Cybersecurity Guidance
Draft June 2013
Final October 2014
Key Principles:
  #1 Shared responsibility between stakeholders, including health care facilities, patients, providers, and manufacturers of medical devices
  #2 Address cybersecurity during the design and development of the medical device
  #3 Establish design inputs for device related to cybersecurity, and establish a cybersecurity vulnerability and management approach as part of the software validation and risk analysis that is required by 21 CFR 820.30(g)
Created electronic cybersecurity review template for review staff to use during premarket review

## Slide 9: Premarket Cybersecurity Submission Expectations
Risk Management (threat modeling)
  Inclusion of hazard analysis, mitigations, and design considerations pertaining to intentional and unintentional cybersecurity risks associated with the device, including:
    A specific list of all cybersecurity risks that were considered in the device design;
    A specific list and justification for all cybersecurity controls that were established for the device.
Traceability
  Inclusion of a traceability matrix that links the actual cybersecurity controls to the cybersecurity risks that were considered

## Slide 10: Premarket Cybersecurity Submission Expectations continued
Lifecycle Plans
  Plan for providing validated software updates and patches as needed throughout the lifecycle of the medical device
  A summary describing controls that are in place to assure that the medical device software will maintain its integrity (e.g. remain free of malware) from the point of origin to the point at which that device leaves the control of the manufacturer
Labeling
  Device instructions for use and product specifications related to recommended cybersecurity controls appropriate for the intended use environment (e.g. anti-virus software, use of firewall)

## Slide 11: Key Principles of FDA Draft Postmarket Management of Cybersecurity in Medical Devices
Collaborative approach to information sharing and risk assessment
Articulate manufacturer responsibilities by leveraging existing Quality System Regulation and postmarket authorities
Align with Presidential EOs and NIST Framework
Incentivize the “right” behavior
Risk-based approach to assuring risks to public health are addressed in a timely fashion

## Slide 12: Core Tenets of Postmarket Management of Medical Device Cybersecurity
Proactively practice good cyber hygiene
Remediate cybersecurity vulnerabilities to reduce the risk to an acceptable level
Conduct appropriate software validation
Properly document the methods and controls used in the design, manufacture, packaging, labeling, storage, installation, and servicing of all finished devices
Identify and implement work-arounds or a temporary fix to adequately mitigate the cybervulnerability risk, even when an “official fix” may not be feasible or immediately practicable
Provide users with relevant information
Emphasis on continuous quality improvement
Emphasis on coordinated disclosure of vulnerabilities and timely response –essential role for stakeholders within the ecosystem to work together

## Slide 13: Use of NIST Framework
Both Guidance documents recommend use of NIST Cybersecurity Framework’s 5 core functions
  Identify
  Protect and Detect
    Vulnerability assessment and risk analysis
  Respond and Recover
    Compensating controls, risk mitigation and remediation

## Slide 14: Cybersecurity – Assessing Risk
Assessment of impact of vulnerability on essential performance and safety of the medical device based on:
Severity of Patient Harm (if the vulnerability were to be exploited)
Exploitability

## Slide 15: Postmarket Cybersecurity Risk Assessment

## Slide 16: Cybersecurity – Assessing Exploitability
Example of elements incorporated into one cyber-vulnerability scoring system:
Attack Vector (physical, local, adjacent, network)
Attack Complexity (high, low)
Privileges Required (none, low, high)
User Interaction (none, required)
Confidentiality (high, low, none)
Integrity (none, low, high)
Availability (high, low, none)
Exploit Code Maturity (high, functional, proof-of-concept, unproven)
Remediation Level (unavailable, work-around, temporary fix, official fix)
Report Confidence (confirmed, reasonable, unknown)
Adopted from: Common Vulnerability Scoring System, V3: Specification Document, available at: www.first.org/cvss/specification-document.

## Slide 17: Controlled Vulnerabilities “Acceptable Residual Risk”
Promote good cyber hygiene and reduce cybersecurity risks even when residual risk is acceptable
Changes to a device solely to strengthen the cybersecurity associated with vulnerability with controlled risk are typically considered device enhancements and/or cybersecurity routine updates and patches, and are not required to be reported
Annual reporting requirements for premarket approval (PMA) devices

## Slide 18: Uncontrolled Vulnerabilities “Unacceptable Residual Risk”
Guidance Addresses:
Reporting Requirements
Time Frame for Mitigating Risks
Public Disclosure
Information Sharing and Stakeholder Collaboration
Notes:     *** The customer communication should, at minimum:
      Describe the vulnerability including an impact assessment based on the manufacturer’s current understanding,
      State that manufacturer’s efforts are underway to address the risk of patient harm as expeditiously as possible,
      Describe compensating controls, if any, and
      State that the manufacturer is working to fix the vulnerability, or provide a defense-in-depth strategy to reduce the probability of exploit and/or severity of harm, and will communicate regarding the availability of a fix in the future,

## Slide 19: Uncontrolled Vulnerabilities “Unacceptable Residual Risk”
Reporting Requirements:
  Manufacturers are required to report uncontrolled vulnerabilities to FDA (21 CFR 806)
  FDA does not intend to enforce reporting requirements under CFR 806 if all of the following circumstances are met:
    No known serious adverse events or deaths associated with the vulnerability
    Within 30 days of learning of the vulnerability, the manufacturer notifies its customers, identifies interim compensating controls, and provides mitigations to bring the residual risk to an acceptable level.
    The manufacturer actively participates as a member of an ISAO.
The manufacturer should evaluate the device changes to assess the need to submit a premarket submission (e.g., PMA, 510(k), etc.) to the FDA
  Remediation of devices with annual reporting requirements (e.g., class III devices) should be included in the PMA annual report, as indicated for controlled vulnerabilities

## Slide 20: Information Sharing and Analysis Organizations (ISAO)
The ISAO best practice models are intended to be:
Inclusive - groups from any and all sectors, both non-profit and for-profit, expert or novice, should be able to participate in an ISAO;
Actionable - groups will receive useful and practical cybersecurity risk, threat indicator, and incident information via automated, real-time mechanisms if they choose to participate in an ISAO;
Transparent - groups interested in an ISAO model will have adequate understanding of how that model operates and if it meets their needs; and
Trusted - participants in an ISAO can request that their information be treated as Protected Critical Infrastructure Information. Such information is shielded from any release otherwise required by the Freedom of Information Act or State Sunshine Laws and is exempt from regulatory use and civil litigation.
See: http://www.dhs.gov/isao

## Slide 21: Key Take Home Messages - for Manufacturers
Design & Develop devices that are securable throughout their product lifecycle
Be mindful that there is an active adversary and that the device will need to be updated so that it can be secure
Software updates for cybersecurity do not require pre-market review or recall (there are some exceptions)
Understand & develop threat modeling for your device
Understand the implications of your own supply chain
Establish a Cybersecurity Risk Management Program
Make cyber hygiene paramount
Respond to and address security vulnerabilities that are identified for your marketed devices
Vulnerability disclosure policy, coordinated disclosure and proactive vulnerability management are critical to improving the security posture of the ecosystem as a whole.
Change the culture of engagement with all stakeholders

## Slide 22: Key Take Home Messages – for Healthcare Delivery Organizations (HDO’s)
Understand what you are purchasing and deploying
Where feasible, include securability for the lifetime of your device in your procurement specs contract language
Develop plan to work with your manufacturers and end users to meet your identified needs
Educate and train your end users on the importance of maintaining system security
Make cyber hygiene paramount
Monitor your network and respond to security vulnerabilities and exploits
Vulnerability disclosure policy, coordinated disclosure and proactive vulnerability management are critical to improving the security posture of the ecosystem as a whole.
Change the culture of engagement with all stakeholders

## Slide 23: Key Take Home Messages – for Security Researchers
Engaging in good faith research towards promoting security and reducing risk of potential harm is very important to the medical device ecosystem
Your technical expertise is of great value and should be leveraged
Be proactive about gaining a better understanding and education of the clinical environment and the regulatory, risk-based framework
Broad assumptions, perceptions and/or entrenched beliefs that stakeholders in healthcare are ignoring the researcher community and have known about these issues for years are just that
Vulnerability disclosure policy, coordinated disclosure and proactive vulnerability management are critical to improving the security posture of the ecosystem as a whole
Change the culture of engagement with all stakeholders

## Slide 24: Summary
Medical device cybersecurity requires a total product life cycle approach: from design to obsolescence
FDA’s proposed regulatory policy incentivizes proactive behavior and good cyber hygiene
Strengthening cybersecurity within the healthcare and public health sector is a collective effort amongst all stakeholders
Development and validation of meaningful tools for assessment of vulnerabilities in the clinical environment is an area of focus going forward

## Slide 25: Next Steps
  Revise DRAFT and release FINAL Guidance
  Continue to promote the development and use of ISAOs within the medical device ecosystem and HPH sector
  Continue to foster collaboration within medical device ecosystem to encourage and support increased adoption of vulnerability disclosure policy and coordinated disclosure
  Leverage positive examples of coordinated disclosure as models for multi-stakeholder engagement
  Continue to build partnerships across other sectors of critical infrastructure and government partners for lessons learned